## Slide 1
KONGSBERG EK80 · WIDEBAND SCIENTIFIC ECHO SOUNDER
Calibrating the EK80
Setup, calibration, data collection and troubleshooting
01
Setup
02
Calibration
03
Data Collection
04
Troubleshooting

## Slide 2
INTRODUCTION
Why the EK80 has to be calibrated
The EK80 measures echo strength electrically. Calibration is what ties that electrical reading to a physical, absolute unit — so a target strength or backscattering value means the same thing on every vessel, on every survey, at every frequency.
1
Accuracy
Converts raw electrical gain into true target strength and volume backscattering — the basis for any quantitative estimate.
2
Comparability
Lets biomass and abundance estimates be compared across vessels, years and survey programmes.
3
Drift & ageing
Transducers and electronics change over time. Recalibration catches gain drift before it biases a survey.
4
Traceability
Produces a documented, dated calibration file that becomes part of the survey's metadata record.
02

## Slide 3
INTRODUCTION
The standard-sphere calibration method
A reference sphere of precisely known target strength is lowered into the beam and moved through it. The EK80 compares its measured echo to the sphere's theoretical TS — any difference becomes a correction.
Transducer
Calibration sphere
known TS
measured echo ↔ theoretical TS
Sphere material
Copper or tungsten carbide, chosen per frequency; each diameter is selected to minimise temperature dependence.
Physical movement
The sphere is deliberately moved through the beam so the wizard can sample both the acoustic axis and off-axis positions.
Output
A calibration file (gain, beam angles, offsets) that the EK80 applies to every subsequent measurement on that channel.
03

## Slide 4
SETUP
01 · SETUP
Choosing a site and casting the environment
Calibration quality depends heavily on conditions. A poor site or missing environmental data will show up later as noisy, low-confidence results — or force the whole protocol to be repeated.
Site checklist
1
Sheltered water
Anchor or hold station somewhere protected from swell and wind chop.
2
Suitable depth
Typically around 50–70 m — enough room to suspend the sphere clear of the surface and seabed.
3
Calm conditions
Low wind and low current keep the sphere still and on-axis in the beam.
4
Daylight, if possible
Reduces the chance of diel vertical migration bringing biological scatterers into the beam.
CTD
Cast a CTD
A temperature/salinity/depth profile at the calibration site feeds directly into the EK80's environment settings.
WHAT IT SETS
Sound speed (c)
used for every range calculation
Absorption coefficient (α)
used for transmission-loss correction
Sphere TS
the reference TS is itself temperature-dependent
04

## Slide 5
SETUP
01 · SETUP
Rigging the calibration sphere
Three monofilament lines, each on its own winch or rod, let the sphere be steered independently in two axes beneath the transducer — and held steady once it's found.
Vessel hull
Transducer
3 winches / rods
Sphere
15–20 m
stabilising weight
1
Loop each line
Tie a secure loop in each monofilament line to hold the sphere.
2
Lower together
Pay out all three lines evenly so the sphere descends roughly on-axis.
3
Mark the lines
Once located, mark each line at its winch — it makes re-locating the sphere far faster next time.
05

## Slide 6
SETUP
01 · SETUP
Preparing the EK80 software
Before the sphere ever goes in the water, the software has to be told what to do — and the channel being calibrated has to be isolated from every other transducer on the vessel.
1
Power up
Turn on the transceivers, then the EK80 display and processing unit, and open the EK80 software.
2
Isolate the channel
Set the channel being calibrated to Active; switch every other channel to Passive so it cannot interfere.
3
Match survey settings
Set transmit power, pulse length and pulse type (CW or FM) to exactly what will be used during the real survey.
4
Set up recording
Configure the raw-data output folder and file settings — always keep the raw data from a calibration.
›
›
›
Why isolate the channel? A transducer left pinging nearby can put stray energy into the one being calibrated, corrupting the sphere echo.
06

## Slide 7
CALIBRATION
02 · CALIBRATION
Finding the sphere in the beam
A weak, off-centre echo almost always means the sphere is sitting in a side lobe, not the main lobe. Small, patient adjustments on each line bring it into the strongest part of the beam.
weak / off-axis
strong / on-axis
Set the display to Sv (20 log) to make the sphere easier to spot against the background.
1
Watch the intensity
Pay each line in and out a little; the echo should get stronger as the sphere nears the axis.
2
Expect some trial and error
It's common to find the weight or sphere sitting in a side lobe at first — that's normal, not a fault.
3
Be patient
Small movements are more effective than large ones once the echo starts to strengthen.
07

## Slide 8
CALIBRATION
02 · CALIBRATION
Mapping the beam pattern
Once located, the sphere is guided through a grid of positions spanning the main lobe. The EK80 calibration wizard logs an echo at each position and fits them to the transducer's beam pattern.
Beam sampled by sphere position
on-axis
Each dot: one sphere position sampled during the pass
FROM THE BEAM MAP, THE WIZARD DERIVES
Transducer gain
on-axis sensitivity, in dB
Beam angles
athwartship / alongship opening angles
Acoustic axis offset
how far the true axis sits from mechanical centre
Sa correction
a small integration correction applied to Sv
Multiple passes through the beam improve confidence in the fit.
08

## Slide 9
CALIBRATION
02 · CALIBRATION
Reading the results — and knowing when to stop
The calibration wizard reports both the corrected parameters and a quality figure. A good calibration is one where that quality figure is tight, not just where the numbers look plausible.
Key outputs
Transducer gain (dB)
On-axis sensitivity — compare against the previous calibration for drift.
Sa correction (dB)
Small correction folded into Sv/backscatter calculations.
Beam angles (dB)
Athwartship and alongship two-way beam widths.
RMS / std. deviation
Scatter of the fitted sphere positions around the model beam — the headline quality metric.
QUALITY RULE OF THUMB
std. dev. ≤ 0.4 dB
Widely used as the threshold for an acceptable sphere-TS standard deviation. Values above this — especially alongside odd beam angles or gain — point to a problem worth investigating before accepting the result.
If in doubt, don't accept it
A poor calibration is worse than a late one — it can be easier and cheaper to repeat the protocol than to correct biased survey data afterward.
09

## Slide 10
DATA COLLECTION
03 · DATA COLLECTION
From calibration file to survey data
A calibration only pays off if the survey actually uses it — with the same acoustic settings it was performed under, and with the data properly logged for later analysis.
1
Apply the file
Load the calibration output into the survey configuration for each channel before the survey starts.
2
Keep settings matched
Transmit power, pulse length and pulse type must stay the same as during calibration — changing them invalidates the gain values.
3
Enable raw logging
Record .raw files continuously; raw data preserves everything needed to reprocess later if settings or software change.
4
Note file structure
Consistent file/folder naming and ping intervals make a survey's archive far easier to work with afterward.
10

## Slide 11
DATA COLLECTION
03 · DATA COLLECTION
Keeping an eye on data quality underway
Calibration establishes accuracy at one point in time. Ongoing checks confirm the system is still behaving the way it was calibrated to, for the length of the survey.
1
Watch the echograms
Scan regularly for unexpected noise bands, dropped pings, or a bottom line that looks wrong for the charted depth.
2
Track background noise
Rising noise floor can mask weak targets long before it becomes obvious — check it against the calibration-time baseline.
3
Use consistent references
A stable patch of seabed or a well-known target class can serve as a quick day-to-day sanity check between full calibrations.
4
Log environment updates
Re-cast sound speed and absorption periodically on longer surveys, especially when crossing water masses.
11

## Slide 12
TROUBLESHOOTING
04 · TROUBLESHOOTING
When calibration doesn't go to plan
!
No / weak sphere echo
The sphere is likely sitting in a side lobe rather than the main lobe, tangled, or on the wrong active channel.
Fix: Re-check that only the target channel is Active; gently work each line to find the strongest echo; verify the sphere hasn't snagged.
!
High scatter / poor RMS
Sea state, current, or a swinging sphere are adding noise to the position fit — not a fault with the electronics.
Fix: Wait for calmer conditions; add or reposition the stabilising weight; take more passes through the beam before accepting.
!
Suspicious parameters
Beam angles, axis offset or gain far from the transducer's nominal values, sometimes paired with high RMS.
Fix: Inspect the transducer face for fouling, marine growth or damage; check connectors and cabling before re-attempting.
!
Unexpected noise
Other acoustic systems (sonars, Doppler logs, other echosounders) or electrical interference corrupt the echo.
Fix: Switch off or set other systems to Passive during calibration; check grounding and shielding on suspect cabling.
12

## Slide 13
BEST PRACTICE
CHECKLIST
Before you call it done
✓
Calibrate before the survey — and again during or after long surveys to catch drift.
✓
Use the exact power, pulse length and pulse type the survey will run.
✓
Record and keep the raw calibration data, not just the summary output.
✓
Re-cast sound speed and absorption if conditions have changed since setup.
✓
Confirm RMS/std. deviation is within tolerance before accepting a channel.
✓
Archive the calibration file alongside dated metadata for every channel.
13

## Slide 14
SUMMARY
Calibration in four phases
1
Setup
Choose a sheltered, suitable site; cast the environment; rig the sphere; prepare the software and isolate the channel.
2
Calibration
Locate the sphere on-axis, map the beam with multiple passes, and confirm the result meets a quality threshold.
3
Data Collection
Apply the calibration, hold settings constant, log raw data, and monitor quality throughout the survey.
4
Troubleshooting
Most problems trace back to beam position, sea state, transducer condition, or interference — diagnose before repeating.